## Slide 1
Feedback from Workshop
Towards harmonised implementation of the
ICH M10 Guideline
ICH M10 team

## Slide 2: The roadmap
Two EBF internal meetings (Zoom) in September/October
  We selected items for discussion this week considering EBF survey and survey to workshop delegates
Face to face Internal EBF members meeting (BCN, 14NOV)
  Deeper dive chapter per chapter which generated draft suggestions for harmonised implementation and identified additional actions
The workshop on 15 Nov to discuss and further calibrate draft suggestions for harmonised implementation or action items

## Slide 3
Next slides:
Key areas of consensus / proposal for harmonised implementation / additional EBF actions
Presented by moderators

## Slide 4
Chapters 1&2
Metabolites in scope
   A good start is to concentrate full validations to late stage molecules e.g. Ph3
Non Clinical PK Studies – examples of what this means would be useful
   When you can’t dose to humans e.g. Anthrax
Method development - What is minimally needed for record keeping?
   Very high level, capture the lifecycle of the assay (if an assay changes after validation capture the why and what has changed)
Why is parallelism not listed as a must have?
   Use scientific judgement – If it adds real scientific value do it, if it doesn’t don’t
Primary vs. secondary matrices – what is primary? What is secondary?
   Comes down to what the data is going to be used for – a good dialogue is required with stakeholders

## Slide 5: Questions – Comments – did we have a blind spot ?

## Slide 6: Chapter 3

## Slide 7: Themes/questions discussed

## Slide 8: (3.2.1 Selectivity)
Majority understand design
Majority support using fewer lots for ‘rare matrices’ (3Rs).
Need EBF definition of ‘rare matrix’ (i.e. small animal species & destructive sampling?)
If selectivity fails, investigate failing batch and/or increase to n=10
(3.2.3 Matrix Effect)
Majority understand design
Majority support using fewer lots/batches for ‘rare matrices’ (3Rs)
Majority agree that we are more likely to receive haemolysed samples rather than lipemic samples (but should assess both haemolysed and lipemic in man)

## Slide 9: (3.2.4 Calibration Curve and Range)
Majority understand design
For a ‘simple’ model, may be sufficient to state linear < quadratic and 1/x < 1/x2 in SOP (depends on dynamic range/instrument)
(3.2.5.1 Preparation of Quality Control Samples)
Majority understand design
MID QC placement (30-50% arithmetic or geometric; your choice)
Majority need clarity on how FAQ links to ICH-M10

## Slide 10: (3.2.5.2 Evaluation of Accuracy and Precision)
Majority understand design and that fresh calibration is used for at least one A&P run
Majority agree that QCs don’t need to be fresh (if frozen, then we need to document stability)
Majority understand design and value of the ‘full size’ robustness run
EBF need to provide examples of A&P calculations (within-day, between day, ANOVA)

## Slide 11: (3.2.7 Dilution Integrity)
Majority support using ‘surrogate matrix’ for ‘rare matrices’ (3Rs)
Majority understand design; 1 bulk, 5 diluted aliquots per DF and 1 replicate per diluted aliquot
Majority interpret 3.2.7 to mean validation bracketing lowest and highest dilution factors (e.g. 1:5 1:10, 1:50, 1:100)
Majority interpret 3.2.7 to mean highest [DQC] ≥ [in-vivo]
Majority understand potential implications if [in-vivo] ≥ [DQC]

## Slide 12: (3.2.8 Stability – Part 1)
Majority interpret 3.2.8 to indicate stability does not need repeating if we have documented evidence (i.e. an approved & traceable BMV-compliant report)
Majority agree stability is performed for each analyte and species
We understand design; 1 bulk, 3 aliquots and 1 replicate (vs independent variables i.e. time and temperature) and including ‘time zero’
Majority agree that frozen run QC may be used (stability proven / will be proven)
Majority agree that stability [DQC] ≥ [in-vivo] (or as high as possible, if [in-vivo] is TBD)
Majority agree that stability of FDC is ‘special’ and does not mean ‘general’ concomitant medication (co-med)

## Slide 13: (3.2.8 Stability – Part 2)
Majority agree that ‘freshly prepared’ means ‘fresh in matrix’ not from freshly prepared primary solutions (SS & WS)
Majority accept design of 3 replicates for Stock And Working Solution Stability (acceptance criteria TBD based on measurement method i.e. UV-VIS, MS or other)
Majority understand Blood Stability is only applicable for plasma assay
Majority accept design of 3 replicates (not aliquots) for Blood Stability (acceptance criteria based on scientific judgment)
Majority agree to limit scope to human Blood Stability (or use human blood as surrogate of animal blood stability; 3Rs)
Recommendation to update EBF position publication on blood stability (Bioanalysis(2011) 3(12) 1333-1336))

## Slide 14: (3.2.9 Reinjection Reproducibility)
Majority agree this’ is different and separate to ‘Processed Sample Stability’ (time between original and reinjection on autosampler)
Majority understand the design of ≥ 5 replicates at LOW, MID and HIGH vs within-run CAL (i.e. reinjection and quantification as ‘self-contained’ run)
Majority agree acceptance criteria of ≤ ±15% accuracy and precision
(3.3.1 Analytical Run)
Majority interpret bracketing to mean at a minimum ‘QC – Samples – QC’?
Majority agree that several ‘batches’ can form a ‘run’ and that QCs (duplicate at three levels) are needed in each ‘batch’
Majority understand carry-over impact and how to mitigate it (note; carry-over only assessed for ‘Run’ not per ‘Batch’)

## Slide 15: (3.3.2 Acceptance Criteria for Analytical Run – Dilution QC)
Majority understand this to mean minimum of ≥ 2 aliquots per DF
Majority accept acceptance criteria; ≥ 50% pass & A @ ≤ ± 15%
Majority agree to bracket the lowest and highest Dilution Factors with DQCs
Majority agree and accept that process control [DQC] ≥ [ULOQ]
Majority agree that if a Dilution Factor (DQC) ‘bracket’ fails, only samples diluted with the passing Dilution Factor (DQC) are accepted

## Slide 16: (3.3.4 Reanalysis of Study Samples)
Majority agree that there are no bioanalytically ‘valid’ study sample concentrations obtained from a failed run
(3.3.3 Calibration Range)
Majority agree and interpret ‘at intended therapeutic dose’ to limit scope of any changes to Phase II/III and Pivotal BA/BE studies

## Slide 17: Questions – Comments – did we have a blind spot ?

## Slide 18: Chapter 4

## Slide 19
Singlicate sample analysis in LBA assays
Reference standard
Critical reagents
QC sample preparation and selection of concentrations
Dilutional linearity
Stability
Selectivity
Surrogate matrix

## Slide 20: Singlicate sample analysis in LBA assays
4.2 Validation
Most often, microtiter plates are used for LBAs and study samples can be analyzed using an assay format of one or more well(s) per sample. The assay format should be specified in the protocol, study plan or SOP. If method development and method validation are performed using one or more well(s) per sample, then study sample analysis should also be performed using one or more well(s) per sample, respectively. ...
EBF position:
European Bioanalysis Forum recommendation on singlicate analysis for ligand binding assays: time for a new mindset
Matthew Barfield, Joanne Goodman, John Hood & Philip Timmerman, Bioanalysis 2020
EBF proposal for implementation
Implement singlicate well analysis as appropriate
If assay validation was performed in duplicate yet data show that singlicate analysis is robust then singlicate sample analysis is justified

## Slide 21: Reference standard
4.1.1. Reference standard
“It is recommended that the manufacturing batch of the reference standard used for the preparation of calibration standards and QCs is derived from the same batch of drug substance as that used for dosing in the nonclinical and clinical studies whenever possible”
EBF proposal for implementation:
If CMC assessments show comparability of drug substance batches, the reference standard is deemed appropriate for bioanalytical purposes

## Slide 22: Critical reagents
4.1.2 Critical Reagents
“Critical reagents, including binding reagents ... have direct impact on the results of the assay and, therefore, their quality should be assured. Critical reagents bind the analyte and, upon interaction, lead to an instrument signal corresponding to the analyte concentration. The critical reagents should be identified and defined in the assay method… .”
“A critical reagent lifecycle management procedure is necessary to ensure consistency between the original and new batches of critical reagents. Reagent performance should be evaluated using the bioanalytical method. ...”
“It can be extended beyond the expiration date from the supplier.”
EBF position and EBF proposal for implementation:
It is positive that these considerations were included in M10
EBF recommendation on practical management of critical reagents for PK ligand-binding assays
  Susanne Pihl, Barry WA van der Strate, Michaela Golob, Laurent Vermet, Birgit Jaitner, Joanne Goodman, Marianne Scheel Fjording & Philip Timmerman, Bioanalysis 2018

## Slide 23: QC sample preparation and selection of concentrations
4.2.4.1 Preparation of Quality Control Samples
The QCs are intended to mimic study samples and should be prepared by spiking matrix with a known quantity of analyte, stored under the conditions anticipated for study samples and analysed to assess the validity of the analytical method.
The analyte should be spiked at the LLOQ, within three times of the LLOQ (low QC), around the geometric mean of the calibration curve range (medium QC), and at least at 75% of the ULOQ (high QC) and at the ULOQ.
EBF proposal for implementation:
Confirm that QCs are spiked correctly before using in formal validation assessment
Use QCs that reflect samples for validation assessment
It is positive that geometric mean is used for the mid QC

## Slide 24: QC sample preparation and selection of concentrations
4.2.4.2 Evaluation of Accuracy and Precision
“If the within-run accuracy or precision criteria are not met in all runs, an overall estimate of within-run accuracy and precision for each QC level should be calculated.“
EBF proposal for implementation:
Follow DeSilva et al (2003)

## Slide 25: QC sample preparation and selection of concentrations
4.3.1 Analytical Run
“The QCs should be placed in the run in such a way that the accuracy and precision of the whole run is ensured taking into account that study samples should always be bracketed by QCs.”
EBF proposal for implementation:
In this instance bracketing means physical placement and not concentration

## Slide 26: Dilutional linearity
4.2.6 Dilution Linearity and Hook Effect
“The same matrix as that of the study sample should be used for preparation of the QCs for dilution.“
“For each dilution factor tested, at least 3 independently prepared dilution series should be performed using the number of replicates that will be used in sample analysis.”
“The dilution factor(s) applied during study sample analysis should be within the range of dilution factors evaluated during validation.”
EBF proposal for implementation:
The use of surrogate matrix may be justified (e.g., free drug) with supporting data
It is positive that it is accepted to use 3 independently prepared dilution series
Dilution QCs do not need to be included in sample analysis if within range of dilution factors tested in validation
Notes: Include an addition slide with conclusions from YEMM discussions?

## Slide 27: Stability
4.2.7 Stability
“Stability evaluations should be carried out to ensure that every step taken during sample preparation, processing and analysis as well as the storage conditions used do not affect the concentration of the analyte.”
EBF proposal for implementation:
Whole blood stability is generally not necessary for large molecules if stability in plasma/serum has been demonstrated under the same conditions unless the analyte is known to behave differently in the presence of blood cells.
  Need for assessment should be driven by modality and mechanism of action.
Stability does not have to be repeated in a second laboratory if conditions and matrix have not changed. The stability report needs to be available and reviewed.

## Slide 28: Stability
4.2.7 Stability
“Since sample dilution may be required for many LBA methods due to a narrow calibration range, the concentrations of the study samples may be consistently higher than the ULOQ of the calibration curve. If this is the case, the concentration of the QCs should be adjusted, considering the applied sample dilution, to represent the actual sample concentration range.”
“For biological drugs, a bracketing approach can be applied, e.g., in the case that the stability has been demonstrated at -70/-80C and at -20C, then it is not necessary to investigate the stability at temperatures in between those two points at which study samples will be stored.”
EBF proposal for implementation:
A single high level stability QC should represent the expected sample concentration (eg close to Cmax).

## Slide 29: Selectivity
4.2.2 Selectivity
“For lipaemic and haemolysed samples, tests can be evaluated once using a single source of matrix. Selectivity should be assessed in samples from relevant patient populations (e.g., renally or hepatically impaired patients, inflammatory or immuno-oncology patients if applicable). In the case of relevant patient populations, there should be at least five individual patients.”
EBF proposal for implementation:
In cases where insufficient samples are available for validation consider in-study selectivity assessment
1 lipaemic and 1 haemolysed sample can be included in the 10 individuals European Bioanalysis Forum: recommendation on dealing with hemolysed and hyperlipidemic matrices. Benno Ingelse, Begona Barroso, Nicholas Gray, Verena Jakob-Rodamer, Clare Kingsley, Corinna Sykora, Petra Vinck, Martina Wein, & Stephen White. Bioanalysis 2014.
Notes: Include an addition slide with conclusions from YEMM discussions?

## Slide 30: Surrogate matrix
2.2.1. Full validation
“In some cases, it may be difficult to obtain an identical matrix to that of the study samples (e.g., rare matrices such as tissue, cerebrospinal fluid, bile or in cases where free drug is measured). In such cases, surrogate matrices may be acceptable for analytical method validation.”
EBF proposal for implementation:
The use of surrogate matrix for calibrators, QCs and sample dilution may be justified with supporting data.

## Slide 31: Questions – Comments – did we have a blind spot ?

## Slide 32: Chapter 5

## Slide 33
Not much is new in ICH M10 related to ISR, see next slide
Recommended to use ICH M10 to reset your compass and not overdo ISR – not all studies are in scope for ISR, they never were and still aren’t
Understand pivotal as where the primary endpoint of the clinical trial is related to PK.
Know your project and understand the purpose of each and every study you support.
Be clear how to distinguish between ISR conduct pr. guidance requirements and how to used ISR to mitigate business risks for pivotal studies.
For guidance on for how, when and where to conduct and document investigation for failed ISR, and for selection of samples for ISR, see Bioanalysis, (2009) 1(6) 1049-1056

## Slide 34: Evolution of ISR
Requirement | ICH M10 | FDA | EMA
Studies in scope | TK species FIH Pivotal BA/BE (Pivotal) FIP (Pivotal) hepatic/renal | All Pivotal for labelling or approval: TK/species All BE, pivotal PK and PD | TK/species BE, FIH, FIP, hepatic/renal
When to execute | Not on the same day as the original analysis | In separate runs | In separate run, at different days
How many samples | 10% first 1000 | 10% first 1000 | 10% first 1000
 | 5% > 1000 | 5% > 1000 | 5% > 1000
Which samples | Near Cmax and elimination phase - representative for the whole study | Near Cmax and elimination phase | Near Cmax and elimination phase
Acceptance Criteria | 2/3 (67%) within 20% (LC/MS), 30% (LBA) | 2/3 (67%) within 20% (LC/MS), 30% (LBA) | 2/3 (67%) within 20% (LC/MS), 30% (LBA)
If it fails… | SOP based investigation | SOP based investigation | Investigation

## Slide 35: Questions – Comments – did we have a blind spot ?

## Slide 36: Chapter 6

## Slide 37
Focus on WHEN and HOW to X-validate
Areas of consensus:
Consultation with stakeholder (clin pharm, regulatory) is key:
  Before X-val occurs:
    BA to explain why X-val is needed
    Clin pharm to explain which studies will be combined
    Decide on what method will be used to assess bias between assays (no acceptance criteria)
  After X-val has occurred
    Clin pharm to assess the impact of any bias between methods
Potential EBF actions:
Is there value in a cybermeeting with stakeholders (e.g. clinical PK/statisticians) on how to “manage” Cross-validation

## Slide 38: Questions – Comments – did we have a blind spot ?

## Slide 39: Chapter 7

## Slide 40
7.1 Methods for Endogenous Compounds
Good to remined  Biomarkers are not in scope.
Modern drugs are often slightly modified versions of endogenous molecules and thus specific bioanalysis is possible: chapter 3 or chapter 4.
The approach adopted (1 of 4 for Chrom, 1 of 2 applicable for LBA) selected based on the needs of the project.
Approach adopted for each program should be scientifically and technology driven.
Further dialogue required in the situation where the dosed compound is not the pharmacologically active entity.

## Slide 41
7.2 Parallelism
You know your drug program, do what makes sense.
When assessed include sufficient samples to understand any limitations.
7.3 Recovery
Consistency of recovery may be assessed in method development.
EBF propose that the well understood extract/post-spike is a valid means of recovery determination.
Most often not relevant for LBA assays: For large molecules recovery may be considered for dried matrix methods/tissue.
7.4 Dried matrix methods
It is welcomed that microsampling is supported to allow consideration of 3Rs and reduce patient burden.

## Slide 42
7.5 Commercial and Diagnostic Kits
Where commercial kits are used, validation should fulfil expectations in section 4 and relevant parts of section 7.
7.6. New or alternative technologies
When technologies not mentioned in the scope of the guideline are used, validation design and acceptance criteria should be science driven.

## Slide 43: Questions – Comments – did we have a blind spot ?

## Slide 44: Chapter 8

## Slide 45
Not recommended as general practice, but it can be a business decision to include a brief history/evolution of the methods in the validation reports when changes have been made to the assay to facilitate the compilation of the CTD modules.
For BA/BE studies, include the additional requested information (e.g. 100% chroms, 100% summary tables of all runs) to the validation reports of methods used to support BA/BE. This information can be added retrospectively by amending the reports ahead of the submission.
Reinjected runs:
  include tables of reinjected runs (including results) to all BA reports.
Rejected runs: for all BA/BE studies, include the tables containing the results from rejected runs and document the reason. Results could be marked as NR in case no reportable data is available.

## Slide 46
SOPs: it was proposed to include the procedures used for the method in the form of a bioanalytical method description (and not to include a list of all SOPs followed to conduct the analysis, as this is information that should be kept at the site).
Instrument I.D.’s: it was recommended to include only the equipment used for data acquisition (i.e. no ancillary equipment).
Sample tracking: it was proposed to indicate only the storage temperature of samples (i.e. no freezer i.d. or location).
Templates (e.g. Summary Tables) would be greatly appreciated (EBF can help)

## Slide 47: Questions – Comments – did we have a blind spot ?

## Slide 48: From here….
Provide FB to global BA community
  meeting report
  Via EBF website these slides
Follow up on actions
  Allow workshop discussion dust to settle and from there, take it further
  Continue to share and evaluate across regions and companies
  EBF commits to include ICH M10 experience discussions in 2023 activities.

## Slide 49
Acknowledgements
49

## Slide 50
Contact Information
Questions: info@e-b-f.eu
European Bioanalysis Forum vzw
www.e-b-f.eu
50